Suggestions for Scenarios for Monthly Safety Drills

- Fire Drill with actual timed egress for patients at pressure evacuating to designated outside safe location (to be performed annually)

- HBO specific fire drill; fire in the chamber, fire in the department, fire in the adjacent area (performed quarterly; refer to operations manual)

- Emergency decompression procedure training for ALL staff and physicians (as required by UHMS; refer to operations manual)

- Confirm the decompression timing of each chamber to and from 3ATA at 5psi/min and emergency vent rate with purge flow rate set at its midpoint (documentation form on NBS website under ‘inspection’, also refer to details listed in Sechrist users manual).

- Verify chamber pressure gauge (refer to Sechrist users manual)

- Verify chamber pressure rate (refer to Sechrist users manual and documentation form on NBS website; ‘inspection’)

- Visually inspect all outdoor chamber exhaust lines for occlusions such as lint and debris. Ensure the exhaust line is covered with a large gauge wire screen no smaller than ½ inch wire mesh per  Sechrist users manual

- Ventilator assembly, operation, and inspection of all parts for signs of wear and tear.

- Nebulizer assembly and operation

- IV pump set up with extension tubing and 3-way stopcock in line for Abbott 3HB and IVAC 530

- Assembly for continuous bladder irrigation in the chamber (refer to NBS policy)

- Blood pressure cuff set up

- Application of cotton Kerlix gauze for hand restraints for in chamber

- Pressure line set up for in-chamber monitoring

- Chamber ground testing

- Meet with local fire department; discuss fire hazards unique to HBO unit and vicinity of oxygen valves, loss of O2 results in immediate decompression of 7psi/min if the FD closes off O2 valves.

- Meet at local fire dept education office to practice using fire extinguisher applying the PASS method

- Transcutaneous oximetry calibration confirmation via local barometric pressure setting, temperature setting at 45C, maximum alarm parameters and zero solution

- Identify and discuss common cardiac dysrhythmias to enhance ECG monitor recognition

- Discuss how to respond to a suicidal patient in chamber is lighter

- Discuss how best to first respond to a patient who describes symptoms of anxiety and/or nausea

- Discuss how best to treat an unconscious CO patient who happens to be a mechanic and smells gas and oil

- Trouble shoot the ventilator that has been setup incorrectly then adjust for multiple settings for each participant

- Describe a plan for how to respond to the intubated patient that loses a secure airway in chamber

- Describe a plan for how to respond to a patient that becomes free of the hand restraints

- Discuss the plan for how to respond to a patient that is locked inside the bathroom and has fallen and cannot get up

- Describe how to respond to the patient who is observed having a seizure at pressure

- Describe how to respond to the patient that complains of sharp right-sided chest pain and pulse of 130

- Describe procedure of how to treat a diabetic with BS of 80

- Discuss and plan how to manage an HBO patient with any of the following:

Wound VAC, Coban dressing, fever, implantable medication pump, defibrillator or pacemaker, external pacemaker, Vaseline gauze dressing, hearing aide, medication patch, diaper, Kotex, Thermacare warmer, On-Q pain pump ball, Unna boot, silver dressing, external fixator, non-gas permeable contact lens, loose fillings, dentures, synthetic hair, pressure ulcers on the coccyx, wedding band unable to be removed, upper respiratory infection, Teed scores >1, anxiety, claustrophobia, steroid use or seizure disorder

Create mock scenarios for the following emergency procedures involving all staff

- Emergency procedure for local oxygen supply failure

- Emergency procedure for hospital wide oxygen supply failure

- Emergency procedure for chamber door safety pin jammed

- Emergency procedure for communication failure

- Emergency procedure for suspected pneumothorax

- Emergency procedure for cardiopulmonary arrest

- Emergency procedure for oxygen toxicity