Concern categories: Energy Source, Fire/Flashpoint, Pressure, Pharmacologic,

Patient Comfort, Acrylic Damage, Infection Control, and Choking Hazard

| MATERIAL | DESCRIPTION | CONCERN | REFERENCE |
| --- | --- | --- | --- |
| BP (NIBP) cuff with tubing | Use of pass-through to connect to Oscillomate 1630 monitor outside chamber. | Energy | NFPA 99 14.2.7.6 |
| Breathing system equipment (built in) | Supply hose with door connection, demand valve mask, mask, face tent, trach collar, non-rebreather masks | Comfort, Pharmacologic | NBS OM |
| Catheter, (indwelling urinary cath or wound drainage system) | Flexible tubing with collapsible reservoir Empty excess drainage, purge air | Comfort, Pressure | NBS P&P |
| Cloth restraints | No velcro or metal components If velcro present TAPE OVER EACH FACE OF VELCRO, if metal present TAPE OVER EXPOSED SURFACES | Energy, Fire, Acrylic Damage | NBS OM P&P |
| Contact lenses (soft and gas permeable) |  | Comfort | SN 0807 |
| Cotton gauze dressings | Pads, packing, wraps, island-type with non-stick surface | Comfort, Pharmacologic | NBS P&P |
| Diaper (adult/infant) /Kotex | Should remain against patient skin, moist, no velcro closures | Comfort, Infection | NBS P&P |
| Dressings (aqueous) or topical medication | Hydrocel | Pharmacologic, Fire, Comfort | NBS P&P |
| ECG and pressure line monitoring cables | Biomed approved/maintained Use of pass-through to connect to monitor outside chamber. | Energy | NFPA 99 14.2.7.6 |
| Eyeglasses | Plastic or metal composition, lenses | Comfort | NBS SN 0903 |
| Feeding tube (Dobhoff) | Flushed and clamped | Comfort, Pressure, Pharmacologic | NBS P&P |
| Gowns, hospital issue | 100% cotton, unit supplied, pocket less, no metal snaps or Velcro, uncontaminated. Must check for concealed items | Comfort, Fire | NFPA 99 14.3.1.5.4 NBS SN 0503 |
| Heimlich Valve | Attached to indwelling chest tube and sterile glove or urinary drainage bag | Pharmacologic, Comfort, Pressure | NBS OM |
| IV extension tubing for Monoplace Hyperbaric chamber | Hospira Use of pass-through to connect to hyperbaric IV pump outside chamber | Pressure | NBS OM P&P |
| Mattress, wedges, pillows | Only those recommended by chamber manufacturer | Comfort, Fire, Pressure | NBS P&P |
| Nasogastric tube | Un-clamped with end loosely covered | Comfort, Pressure | NBS P&P |
| Nebulizer equipment (in-chamber) | Disposable nebulizer, 5-in-1 adapter, liquid medication | Pharmacologic, Comfort | NBS P&P |
| Plastic drink cup | No-spill style, vented, unit supplied, only approved unit supplied fluid | Comfort, Pressure, Fire, Pharmacologic | NBS SN |
| Plastic tape | No paper or silk tape | Fire, Energy | NFPA 99 14.3.1.5.4.1 |
| Pressure infuser with spacer | Ethox Surgi-Press, oxygen-safe lubricant (silicon base) | Pressure, Pharmacologic | NFPA 99 14.3.6.2.1.4 |
| Sheets, pillow cases, towels, wash cloths, blankets | 100% cotton, unit supplied | Comfort, Fire | NFPA 99 14.3.1.5.4.2 |
| TCOM electrode with cord | Radiometer – Max 28v and 4w in class B chamber Use of pass-through to connect to monitor outside chamber | Energy | NFPA 99 14.2.7.6.1, 14.2.7.6.1.1 |
| Urinal, bedpan | Disposable, plastic | Comfort |  |
| Ventilator breathing circuit (hyperbaric) | Sechrist Model 500A Use of gas port to connect to control module outside chamber | Pressure, Pharmacologic | NBS OM P&P |

ReferenceTerminology:

NFPA; National Fire Protection Association 99 2012 Edition

NBS SN; National Baromedical Services Safety Notice

NBS OM; National Baromedical Services Operations Manual

NBS P&P; National Baromedical Services Policy and Procedure Manual